## Slide 1: Optical Characterization setup
Kazuhiro Hotta

## Slide 2: Self-driving laboratory
AI algorisms
Quantum chemical
calculations
Automated
robotic platform
Property predictions
Experiment
Accelerate the materials development by efficiently exploring huge and complex search spaces.
F. Häse et al., Trends in Chemistry 1, 282-291 (2019)

## Slide 3: MADNESS Project
High performance computing
for quantum-based simulations
- property prediction
- reaction condition prediction
AI-aided
experiment
planning
Automated Chemical synthesis system
Automated Characterization
Platform
Define search space (building LEGO blocks)
...
LEGO like synthesis based on Suzuki-Miyaura coupling
Accelerate the materials development by efficiently exploring the huge search space.

## Slide 4: My task
Development of the automated characterization setups for the OLSD molecules
Integrating them with the automated synthesis and identification/separation platforms
Synthesis
Identification/separation
Characterization

## Slide 5
Target properties for laser molecules

## Slide 6: Target
~ Material for the organic semiconductor lasers ~
Advantages over
inorganic lasers
Low-cost fabrication
Solution processability
Mechanical flexibility
Wavelength tunability
Promising
applications
A.S.D. Sandanayaka et al., Appl. Phys. Express 12, 061010 (2019)
Technical difficulties
Next Gen.
Communication
Heads up display
Thermo dynamic therapy
Triplet and polaron related losses
  optical absorption
  single-triplet annihilation
Degradation in the intense excitation
Low carrier mobility
Light amplification in
the Laser cavity
Stimulated emission

## Slide 7: Dynamics in OSLD
Gain
Cavity loss
Absorption losses
Photon
density
Photon
flux
Spontaneous emission
Model of OSLD device architecture
Hole injection
Light confinement
Recombination
Electron injection
Light outcoupling

## Slide 8: Target properties
2. Small ground and excited state absorption at the emission wavelength
Estimated relative populations during lasing
High Gain
(large gain cross section and singlet population)
Absorption cross section should be smaller than 𝝈g in proportion to the relative populations
Fast radiative decay
Narrow emission bandwidth
Large fluorescence quantum yield
Small probability of singlet related annihilations
)
Gain should be larger than the sum of all optical losses
Lasing condition

## Slide 9: Progress in automated characterization setups
UV-VIS absorption/
Photoluminescence(PL)
- Emission band width
Target property :
- PL quantum yield
Development Status : Done
- Ground state Abs.
Photoluminescence lifetime
- Radiative decay constant
Target property :
Development Status :
Triplet absorption
- Triplet absorption
Target property :
Development Status :
- Implementing additional lasers
- Prototyping was done
Done
Spectro-electrochemistry
- Charged state Abs.
Target property :
Development Status :
- Got potentiostat
- Selecting spectrometer

## Slide 10
Optical Characterization setups

## Slide 11: Communication workflow
HPLC-MS
Characterization
setups
HPLC-MS job
HPLC-MS
status
HPLC-MS job
HPLC-MS/Sample collector status
Characterization job
Sample collector
status
Sample collector
status
Characterization job
Sample collector status
Dropbox
Reaction mixture
injection port
Purified reaction product
Chemspeed

## Slide 12: Current setup
UV/VIS
PL lifetime
Inject
Waste
Chemspeed
HPLC-MS
Pump
DAD
Waste
MS
Collection Vials,
Heater/Chiller/Shaker
Waste
Solvent
Syringe pump
Dilution vial
Waste
PL
Column
nitrogen

## Slide 13: Optical setups
1 : Syringe pump
2 : 6-port valve
3 : flow splitter
4 : solvent bottle
5 : waste bottle
6 : UV/VIS absorption cell
7 : PL cell
8 : PL lifetime setup
9 : power meter
10 : spectrometer
11 : D2/Halogen lamp
12 : shutter driver
13 : pico-second laser driver
14 : oscilloscope/function generator
15 : collection vials
16 : shaker/heater/chiller
(evaporation/redissolution unit)
From
HPLC-MS
1
2
3
4
5
6
7
12
9
10
11
8
13
14
15
16
17

## Slide 14: The PL lifetime measurement setup
System response function : ca. 0.2 ns
Fully automated measurement and data analysis
The flow cell is connected to the HPLC-MS
Throughput : < 1 min/sample
 | DSB* | PTPTP* | Rhodamine 6G | Rhodamine B
Our results (ns) | 1.29 ± 0.03 | 0.95 ± 0.03 | 3.95 ± 0.06 | 1.79 ± 0.09
Literature (ns) | 1.26 | 1.03 | 3.87 | 1.92
* The error is 3𝛔 obtained from three measurements
PTPTP
From
HPLC
To
waste
To flow cell
dilution
vial
solvent
Sample
loop
APD
(detector)
ND filter
wheel
Flow cell
Pico second laser
lens
lens
mirror
Inside of the box

## Slide 15: Optical setups ~schematic diagram~
ND filter wheel
UV/VIS
cell
PL lifetime
cell
Solvent
Syringe pump
Dilution
vial
Waste
PL cell
Waste
LED
D2/Halogen
lamp
P.M.
shutter
Spectrometer
optical
fiber
APD
Pico second
Pulse laser
bandpass
filter
Driver
TCSPC
board
Function
generator
From
HPLC-MS
Collection Vials,
Heater/Chiller/Shaker
nitrogen

## Slide 16: Optical setups ~schematic diagram in detail ~
shutter
PC
UV/VIS
cell
PL lifetime
cell
Solvent
Syringe pump
Dilution
vial
Waste
PL cell
LED
D2/Halogen
lamp
P.M.
Spectrometer
optical
fiber
APD
Pico second
Pulse laser
bandpass
filter
Driver
TCSPC
board
Function
generator
From
HPLC-MS
Collection Vials,
Heater/Chiller/Shaker
N2
ND filter wheel
6-port
valve
flow
selector
Valve
driver
Lamp
Controller
(USB GPIO)
LED driver
Shutter
driver
Power meter
console
Waste
Solution line
Electrical
connection

## Slide 17: Automated workflow (common for PL, Abs, TE)
Initial state :
  The sample solution is in the dilution vial
  The flow cell is filled with a solvent (typically ACN)
2. Dark (PL and Abs. ) and Reference (Abs. ) measurement:
  Measure reference spectrum (after dark measurement with shutter close)
3. Sample transfer:
  Send air to the cell to remove solvent from the cell*
  Withdraw sample solution from the dilution vial and send it to the cell
air (or N2)
air (or N2)
4. Sample measurement:
  Measure sample spectrum (after dark measurement with shutter close)
5. Sample collection:
  Collect sample from the cell and send it to the dilution vial for the following measurement**
1. and 2.
3-1.
3-2.
3-2.
4.
5.
5.
* better not to withdraw solvent from the cell to prevent unintentional sample dilution due to the dead volume (~ 20 uL) of the syringe.
** cell washing is done only after all the measurements (PL, TE) finished for the same reason.

## Slide 18
Fiber BF19Y2HS02
(f = 0.8 mm, NA 0.22, q = 12.7o)
24mm
51.6mm

## Slide 19
Automation of the identification, purification and characterization workflow

## Slide 20: Sample identification/separation system based on the HPLC-MS
Waste
UV-VIS
detector
Mass spec
pump
Column
(separation)
Collection for characterization
(or to waste)
Crude reaction
mixture
200
250
300
350
400
Wavelength / nm
0.0
2.5
5.0
7.5
10.0
12.5
15.0
Retention Time / min
absorbance chromatogram
Retention Time / min
m/z
MS chromatogram*
Chemspeed
ThermoFisher HPLC-MS system

## Slide 21: Software for the automation of HPLC-MS
Mass spec reader
/analyzer
Database interface
HPLC-MS control
/data access
Chromatogram analyzer
Modules
ThermoHPLCMS.py
ThermoFisher
Chromeleon 7 SDK
ThermoFisher
RawFileReader
PythoMS*
MADNESS
DB**
Core
m/z
Ion count
Python package for the automated
HPLC-MS experiment
Instrument control
data analysis
interact with other devices
interact with the database
* Yunker et al. J. Chem. Inf. Model. 2019, 59, 4, 1295–1300.

## Slide 22: Automated peak analysis
Collapse 3D
to 2D
200
250
300
350
400
Wavelength / nm
Time / min
3D absorbance chromatogram
Background
correction
Background
Correction/
Peak detection
3D Peak deconvolution
(MCR-ALS*)
200
250
300
350
400
Wavelength / nm
Time / min
component 1 : r. time, peak area, …
component 2 : r. time, peak area, …
component 3 : r. time, peak area, …
:
2D/3D Peak deconvolution
for overlapping peaks
(MCR-ALS*, gaussian, exponentially modified gaussian, ..)
* Multivariate Curve Resolution-Alternating Least Squares
(M. M. Sinanian et al., Anal. Chem. 88, 11092−11099 (2016))
Chromatograms and absorption spectra
for each component

## Slide 23: Peak identification in the targeted measurement
Retention Time / min
m/z
3D MS chromatogram*
monoisotopic mass (Mm) of the target
Selected ion chromatogram (XIC)
at the Mm of the target
Isotope pattern confirmation
for the target
to the characterization setups

## Slide 24: Hardware configuraion
Waste
Inject
Waste
Solvent
6-port
valve
Syringe
pump
Vial for
dilution
Waste
Synthesis
Identification/quantification/separation
Characterization
PL lifetime
Flow cell
PL
Flow cell
Absorption
Flow cell
To optical setups
Chemspeed
injection port
To HPLC
From HPLC
Flow
splitter
Waste
PL lifetime setup
Absorption flow cell
PL flow cell
From
HPLC-MS
Solvent
6-port valve
Flow
selector
Collection vials
Vials
Flow
selector
Syringe
pump

## Slide 25
Inject
Waste
Selected ion chromatogram (XIC)
time / min
time / min
Retention time based multiple targets collection
Target identification based on the XIC
Parallel HPLC-MS and characterization
Waste
Solvent
6-port
valve
Syringe
pump
Vial for
dilution
Waste
Synthesis
Identification/quantification/separation
Characterization
PL lifetime
Flow cell
PL
Flow cell
Absorption
Flow cell
Flow
splitter
Waste
Flow
selector
Collection vials
UV-vis absorption chromatogram

## Slide 26: Characterization results
time / s
time / ns
BSBCz
DSB
BSBCz
DSB
Abs at 365nm (Dashed)
PL at λem
(solid)
Photodegradation with O2*
* For BSBCz, the double bonds are oxidized
and then broken
Absorption and PL spectrum
wavelength / nm
Time resolved PL intensity
Absorbance and PL time trace
Absorption and PL spectrum
Throughput : About 10 minutes/sample including system washing

## Slide 27: Requirement for the OSL molecules
Find the maximum in the region with no absorption
Absorbance/PL spectrum
Requirement for the laser
Gain should be larger than the sum of all optical losses
Gain cross-section
Spectral factor
Absorbance/Calculated spectral factor

## Slide 28: Objective functions for the OSL molecules
*error bar : 3𝛔 of three measurements
Gain cross-section
Spectral factor
(1E-24 cm2 s)
Relative QY
1/PL lifetime
(1E9 s-1)
PL Retention
(at 1 min)
Spectral factor

## Slide 29
Demonstration of the workflow

## Slide 30: One-pot synthesis of BSBCz like molecules

## Slide 31: Molecules to be tested

## Slide 32: Automated synthesis and characterization workflow

## Slide 33: Characterization results
…

## Slide 34: Characterization results

## Slide 35: Characterization results